| Approval and Issue record: |  |  |  |  |
| --- | --- | --- | --- | --- |
| Revision No. | Date | Description (see Revision Record for details) | Author (name) | Approved (name) |
| Revision Record: |  |  |  |  |
| Revision No. | Description of Revision |  |  |  |
| Proposed Revision by document user: |  |  |  |  |
| Name of Proposer | Revision Suggested (please detail line numbers) |  |  |  |
| Document References relating to equipment/system: |  |  |  |  |
| Document Number | Document Title |  |  |  |
| 1200-MOG60-OPPS-0010-001 | Permit To Work Management |  |  |  |
| 1200-MOG60-OPPS-0005-001 | Maintenance Execution |  |  |  |
| 4103-50-6700-G05-0001-001 | Installation, Operation & Maintenance Manual (IOM) |  |  |  |
| 4103-50-6700-G23-0001-001 | FAT / Performance / Functional Test Procedures Metering Control System |  |  |  |
| 4103-50-6700- G23-0002-001 | FAT / Performance / Functional Test Procedures Metering Control System |  |  |  |
| 4103-50-6700-G23-0003-001 | FAT / Performance / Functional Test Procedures Metering Control System |  |  |  |
| 4103-50-6700- G23-0004-001 | FAT / Performance / Functional Test Procedures Metering Control System |  |  |  |
| 4103-50-6700- G23-0005-001 | FAT / Performance / Functional Test Procedures Metering Control System |  |  |  |
| 4103-50-6700- G23-0006-001 | FAT / Performance / Functional Test Procedures Metering Control System |  |  |  |
| 4103-50-6700- G11-0001-001 | Manufacturer’s Catalogue – Turbine Meters |  |  |  |
| 4103-50-6700- G11-0003-001 | Manufacturer’s Catalogue - Transmitters |  |  |  |
| 4103-50-6700- G11-0004-001 | Manufacturer’s Catalogue – Density/Viscosity Analyzer |  |  |  |
| 4103-50-6700- G11-0005-001 | Manufacturer’s Catalogue – BS&W Analyzer |  |  |  |
| 4103-50-6700- G11-0006-001 | Manufacturer’s Catalogue – Grab Sampler |  |  |  |
| 4103-50-6700- G11-0007-001 | Manufacturer’s Catalogue - FCV’s |  |  |  |
| 4103-50-6700- G11-0009-001 | Manufacturer’s Catalogue – Flow Computers |  |  |  |
| 4103-MI20-50PO-9937-001 | Gas Chromatograph – Requisition Specification |  |  |  |
| 2500-M380-50S1-9937-001 | GAS CHROMATOGRAPH - STANDARD SPECIFICATION |  |  |  |
| IEC 61508/61511 | IEC Safety Integrity Levels (SIL) |  |  |  |
| IEC 60079-10-1 | Hazardous Area Classification |  |  |  |
| ISO 14001 | Environmental Standards |  |  |  |
| ISO 3171 | Petroleum liquids — Automatic pipeline sampling |  |  |  |
| API MPMS Chapter 4, 5, 8 & 10 | American Petroleum Institute Manual of Petroleum Measurement Standards. |  |  |  |
| 2500-M380-50DA-0001-001 | INSTRUMENT INSTALLATION STANDARDS |  |  |  |
| 2500-M380-50DA-0002 | PROCESS AND PNUEMATIC HOOK UP STANDARDS |  |  |  |
| CAL-1211 | CAL-1211_EN - CHROMATOGRAPH ANALYZER CALIBRATION |  |  |  |
| CAL-1209 | CAL-1209_EN - BSW ANALYZER CALIBRATION |  |  |  |
| CAL-1210 | CAL-1210_EN - DENSITY ANALYZER CALIBRATION |  |  |  |

| Task Details: |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Purpose of Assessment |  |  | The purpose of this assessment is to evaluate the competence of personnel working on oil and gas samplers and analyzers aboard the FPSO UARA, ensuring their ability to safely, accurately, and efficiently operate, maintain, and troubleshoot sampling and analytical systems. These systems are critical for quality assurance, custody transfer, regulatory compliance, and process optimization. |  |  |  |  |  |  |  |  |  |  |  |
| Process / Equipment |  |  |  |  |  | Task Name |  |  |  |  |  |  |  |  |
| Date |  |  |  |  |  | Reassessment? |  |  | Yes ☐ |  |  | No ☐ |  |  |
| Candidate |  |  |  |  |  | Assessor |  |  |  |  |  |  |  |  |
| Name |  |  |  |  |  | Name |  |  |  |  |  |  |  |  |
| Job Title |  |  |  |  |  | Job Title |  |  |  |  |  |  |  |  |
| Asset |  |  |  |  |  | Asset |  |  |  |  |  |  |  |  |
| Signature |  |  |  |  |  | Signature |  |  |  |  |  |  |  |  |
| Assessment Outcomes |  |  |  |  |  |  |  |  |  |  |  |  |  |  |
| Assessment Evidence Requirements: |  |  |  |  |  |  |  |  |  |  | Yes |  | No | N/A |
| 1 | Has the individual been observed carrying out the task? |  |  |  |  |  |  |  |  |  | ☐ |  | ☐ | ☐ |
| 2 | Does the evidence show ALL skills requirements are met? |  |  |  |  |  |  |  |  |  | ☐ |  | ☐ | ☐ |
| 3 | Does the evidence show ALL knowledge requirements are met? |  |  |  |  |  |  |  |  |  | ☐ |  | ☐ | ☐ |
| 4 | Has candidate provided sufficient valid, authentic, reliable, and current evidence? |  |  |  |  |  |  |  |  |  | ☐ |  | ☐ | ☐ |
| Proficiency Level Assessed |  |  |  |  | Basic ☐ |  | Intermediate ☐ |  |  | Advanced ☐ |  |  |  |  |
| Assessment Outcome |  |  |  |  | Competent ☐ |  |  |  |  | Not Yet Competent ☐ |  |  |  |  |
| Training needs identified? |  |  |  | Details of Training needs: |  |  |  | Candidate Comments on Assessment |  |  |  |  |  |  |
| Yes ☐ |  | No ☐ |  |  |  |  |  |  |  |  |  |  |  |  |
| Assessor Overall Comment |  |  |  |  |  |  |  |  |  |  |  |  |  |  |

| Assessment Planning / Candidate Briefing |  |  |
| --- | --- | --- |
| Evidence to be Collected | - Records of outputs/ actions arising from observation (primary method) - Written or verbal responses to questions - Evidence of prior work, procedural development - Records of outputs/ actions arising from demonstrations and/ or simulations. |  |
| Candidate to answer and tick the following questions: |  |  |
| Has the purpose and consequences of the assessment been explained? |  | ☐ |
| Has the assessment been planned and given me sufficient time to prepare? |  | ☐ |
| Have the assessment methods been agreed and explained? |  | ☐ |
| Do you understand what evidence is to be collected and why? |  | ☐ |
| Do you understand the assessment procedure and the preparation you should undertake? |  | ☐ |
| Have your rights and the appeal system been explained? |  | ☐ |
| Have you discussed any special needs or conditions to be considered during the assessment? |  | ☐ |
| Are you willing and ready to be assessed? |  | ☐ |
| Have you agreed on a suitable time and place for the assessment? |  | ☐ |
| Do you have any objections to or concerns about the proposed assessor? |  | ☐ |
| The assessment methods for the assessment have been agreed and I understand them? |  | ☐ |
| Due to the availability of process equipment and production requirements, some tasks may require to be simulated as part of the assessment. |  | ☐ |
| The task will be performed in the normal manner and should only be interrupted by the Assessor if safety or production is compromised. |  | ☐ |

| Assessor is to check the relevant methods used: 1. Observed 2. Simulated/Demonstrated 3. Questioned |  | 4. Professional Discussion 5. Witness Testimony 6. Other (Please comment on methods used) | Check Relevant Boxes |
| --- | --- | --- | --- |
| Code | Underpinning Knowledge |  |  |
| Basic – Level 1 |  |  |  |
| K. 1.0 | Explain the principles of oil and gas metering, including fiscal measurement and accuracy standards. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 1.0 | - Verify understanding of custody transfer and its reliance on accurate measurement. - Confirm knowledge of industry standards such as API MPMS Chapters 4, 5, and 8. - Assess understanding of measurement components: flow, mass, density, water cut, and viscosity. - Ensure awareness of accuracy requirements (±0.1% flow, ±0.01% proving repeatability). |  |  |
| K. 1.1 | Describe the purpose and operation of a prover loop in metering calibration. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 1.1 | - Verify understanding of bi-directional sphere provers for validating flowmeter accuracy. - Confirm knowledge of calibrated volume repeatability (±0.01% mean volume). - Assess awareness of flowrate limits and maximum sphere velocity (1.6 m/s). - Ensure understanding of pressure drop management during proving (<2.2 bar). - Verify the candidates understanding of the maintenance procedures and work orders. |  |  |
| K. 1.2 | Explain the principle of differential pressure measurement for flow. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 1.2 | - Verify understanding of Bernoulli’s Principle and its application in DP transmitters. - Confirm knowledge of impulse tubing setup, static pressure, and span calibration. - Assess awareness of range limits (e.g., 0–250 mbar) and flow accuracy requirements. - Ensure understanding of pressure and temperature compensation in flow calculations. |  |  |
| K. 1.3 | Describe the setup and configuration of HART communication for DP transmitters. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 1.3 | - Verify ability to connect HART communicators safely to the transmitter. - Confirm understanding of accessing primary variables, diagnostics, and configuration menus. - Assess knowledge of zero/span calibration and setting upper/lower range values. - Ensure understanding of error code troubleshooting via HART communication. |  |  |
| K. 1.4 | Describe the operational principles of a helical turbine meter. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 1.4 | - Verify understanding that flow velocity determines turbine rotation, producing pulses proportional to flowrate. - Confirm awareness of flow conditioning requirements (10D/5D straight runs). - Assess understanding of meter accuracy limits (±0.5% steady flow). - Ensure awareness of pressure drop and flowrate limitations. |  |  |
| K. 1.5 | State the maintenance requirements for a Heliflu-TZN turbine meter. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 1.5 | - Verify awareness of periodic cleaning and inspection of the turbine and bearings. - Confirm knowledge of checking for wear or debris affecting turbine performance. - Assess understanding of monitoring pulse outputs for consistency. - Ensure familiarity with lubrication requirements and calibration schedules. |  |  |
| K. 1.6 | Explain the operating principle of the Zelen Tech ZT-100 water cut analyzer. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 1.6 | - Verify understanding of capacitance-based water detection and dielectric constant measurement. - Confirm knowledge of operational range (0.01%–1.0% v/v). - Assess awareness of temperature and flow effects on water cut readings. - Ensure understanding of calibration requirements using water-in-oil standards. |  |  |
| K. 1.7 | List the common faults and troubleshooting methods for the ZT-100 analyzer. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 1.7 | - Verify ability to identify sensor fouling, flow inconsistencies, or electrical faults. - Confirm knowledge of validating sensor outputs using manual sampling. - Assess understanding of recalibrating probes after drift detection. - Ensure ability to use diagnostics to isolate signal loss or noise. |  |  |
| K. 1.8 | Explain the operating principle of the Emerson Micro Motion viscosity transmitter. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 1.8 | - Verify understanding that viscosity is measured using torsional oscillation of a vibrating tube and its response to fluid properties. - Confirm knowledge of the standard viscosity units: centipoise (cP) or centistokes (cSt). - Assess awareness of the role of temperature compensation to correct viscosity values. - Ensure understanding of the relationship between viscosity, temperature, and flow dynamics. |  |  |
| K. 1.9 | State the maintenance requirements and common faults for the viscosity transmitter. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 1.9 | - Verify awareness of sensor fouling due to heavy hydrocarbons or contamination. - Confirm knowledge of periodic cleaning and checking transmitter diagnostics. - Assess understanding of signal drift and recalibration requirements. - Ensure familiarity with fault indicators such as low amplitude, frequency errors, or transmitter alarms. |  |  |
| K. 1.10 | Describe the purpose and key functions of the ABB Spirit Flow-X Flow Computer. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 1.10 | - Verify understanding that the Flow-X is used for totalizing flow, density, and BS&W measurements in fiscal metering. - Confirm awareness of its digital and analog inputs/outputs for instrument communication. - Assess knowledge of real-time data logging, trend analysis, and alarm management capabilities. - Ensure familiarity with communication protocols such as Modbus, HART, or PROFIBUS. |  |  |
| K. 1.11 | Explain the role of alarms, error logs, and diagnostics in the Flow-X system. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 1.11 | - Verify understanding of critical alarms (e.g., high flowrate, low differential pressure) versus non-critical alarms. - Confirm awareness of how to access and interpret fault logs and system diagnostics. - Assess ability to identify issues like signal loss, sensor errors, or communication faults. - Ensure understanding of documenting and escalating unresolved alarms to supervisors. |  |  |
| K. 1.12 | Explain the purpose and operation of the Auto Control CS55 grab sampler. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 1.12 | - Verify understanding of flow-proportional sampling to collect representative oil samples for analysis. - Confirm awareness of the role of grab sampling in validating BS&W analyzers and metering accuracy. - Assess knowledge of the mechanical operation, including sample timing and valve actuation. - Ensure understanding of proper handling and preservation of samples to maintain integrity. |  |  |
| K. 1.13 | State the safety and operational checks required for grab sampling. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 1.13 | - Verify understanding of isolating the system before opening sample valves. - Confirm knowledge of ensuring sample points are flushed to avoid contamination. - Assess awareness of handling pressurized systems and preventing spills. - Ensure understanding of safe storage, labeling, and transportation of collected samples. |  |  |
| K. 1.14 | Explain the process for documenting and escalating metering discrepancies to supervisors. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 1.14 | - Verify understanding of identifying deviations in flowrate, BS&W, or density during operations. - Confirm knowledge of compiling findings in daily/shift reports with accurate timestamps and meter readings. - Assess ability to explain the importance of escalation for unresolved issues or critical alarms. - Ensure understanding of adhering to reporting templates and organizational standards. |  |  |
| K. 1.15 | Describe the importance of diagnostic data when reporting equipment faults. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 1.15 | - Verify ability to access and interpret diagnostic logs (e.g., ABB Flow-X, HART diagnostics). - Confirm understanding of documenting errors such as signal drift, sensor faults, or hardware failures. - Assess knowledge of reporting fault details with error codes, observed trends, and attempted corrective actions. - Ensure understanding of highlighting faults that may require calibration, repair, or replacement. |  |  |
| K. 1.16 | Explain the working principle of a Gas Chromatograph. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 1.16 | - Uses carrier gas to transport the sample through a chromatographic column. - Separates gas components based on boiling points and adsorption properties. - Detects and quantifies gas components using a Thermal Conductivity Detector (TCD) or Flame Ionization Detector (FID). - Outputs gas composition data to the control system for process monitoring. |  |  |
| K. 1.17 | Identify the main components of a Gas Chromatograph system. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 1.17 | - Sample probe extracts a representative gas sample from the process stream. - Sample conditioning system (SCS) regulates pressure, removes contaminants, and ensures proper gas flow. - Chromatographic column separates gas components for individual detection. - Carrier gas supply system maintains stable transport of gas through the system. |  |  |
| Intermediate -Level 2 |  |  |  |
| K. 2.0 | Explain the principles and standards of oil and gas sampling (ISO 3171, API MPMS Chapter 8). |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 2.0 | - Verify understanding of flow-proportional sampling techniques for crude oil and gas streams. - Confirm knowledge of standards ensuring representative sampling under varying flow conditions. - Assess awareness of sample conditioning systems (e.g., fast loops, inline heaters). - Ensure understanding of sampling frequency, sample preservation, and contamination prevention. |  |  |
| K. 2.1 | Describe the operation and purpose of automatic grab samplers (Auto Control CS55). |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 2.1 | - Verify knowledge of flow-proportional and time-based sampling mechanisms. - Confirm understanding of isolating pressurized lines for safe sample collection. - Assess ability to explain maintenance requirements such as cleaning valves and seals. - Ensure familiarity with sample preservation methods for accurate analysis. |  |  |
| K. 2.2 | Explain the working principles of water-in-oil monitors (Zelen Tech ZT-100). |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 2.2 | - Verify understanding of capacitance-based measurement and dielectric constants. - Confirm knowledge of operational ranges (e.g., 0.01–1.0% v/v) and calibration techniques. - Assess awareness of temperature and flow impacts on analyzer readings. - Ensure ability to troubleshoot signal drift caused by fouling or probe degradation. |  |  |
| K. 2.3 | State the calibration and maintenance requirements of density analyzers (e.g., Emerson Micro Motion). |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 2.3 | - Verify understanding of zero/span checks and use of certified reference fluids. - Confirm knowledge of fouling impacts on sensor performance and signal quality. - Assess understanding of pressure and temperature compensation for accurate readings. - Ensure awareness of industry standards for density measurement accuracy. |  |  |
| K. 2.4 | Explain the principles of gas chromatography and its role in gas composition analysis. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 2.4 | - Verify understanding of column separation and detector analysis methods. - Confirm knowledge of calibration using gas standards for accurate composition. - Assess awareness of methane, H2S, and heavier hydrocarbon measurement requirements. - Ensure understanding of reporting results for flare gas management or custody transfer. |  |  |
| K. 2.5 | Describe the purpose of viscosity measurement in crude oil sampling (Micro Motion Transmitters). |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 2.5 | - Verify knowledge of torsional oscillation-based viscosity measurement techniques. - Confirm understanding of viscosity units (cP, cSt) and their process impacts. - Assess ability to describe temperature compensation and fault diagnostics. - Ensure awareness of calibration intervals for maintaining measurement accuracy. |  |  |
| K. 2.6 | State the safety protocols for hazardous area maintenance of analyzers (ATEX, IECEx). |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 2.6 | - Verify understanding of isolating analyzers before maintenance. - Confirm knowledge of gas testing and safe working practices in Zone 1/Zone 2 areas. - Assess awareness of maintaining electrical equipment certifications in hazardous locations. - Ensure understanding of PPE, permit systems, and lockout/tagout (LOTO) procedures. |  |  |
| K. 2.7 | Explain the purpose of regulatory compliance for sampling and analysis activities. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 2.7 | - Verify knowledge of meeting environmental standards for emissions monitoring and water quality. - Confirm awareness of audit trail documentation, including calibration and maintenance logs. - Assess understanding of reporting processes for custody transfer and environmental compliance. - Ensure familiarity with ISO, API, and local regulatory requirements. |  |  |
| K. 2.8 | Describe techniques for supervising and optimizing sampling campaigns. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 2.8 | - Verify understanding of planning sampling schedules based on production and regulatory needs. - Confirm knowledge of resource allocation (technicians, tools, spare parts). - Assess awareness of ensuring representative samples under fluctuating process conditions. - Ensure familiarity with reporting sampling campaign results and identifying optimization areas. |  |  |
| K. 2.9 | Explain the importance of sample conditioning in Gas Chromatograph operation. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 2.9 | - Removes liquid hydrocarbons, water, and particulates that can damage the column. - Ensures consistent sample pressure and flow rate for accurate measurement. - Uses filters, pressure regulators, and bypass loops for system protection. - Prevents contamination-related measurement drift and response delays. |  |  |
| K. 2.10 | Analyze calibration and validation requirements for a Gas Chromatograph. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 2.10 | - Uses certified calibration gases (CGA 580 for carrier gas, CGA 660 for calibration gas). - Performs baseline calibration checks before routine operation. - Conducts auto-calibration cycles at scheduled intervals. - Ensures compliance with ASTM D3764-19 validation practices. |  |  |
| Advanced -Level 3 |  |  |  |
| K. 3.0 | Explain the principles of custody transfer metering and its role in fiscal reporting. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG.3.0 | - Verify understanding of custody transfer metering accuracy requirements (e.g., ±0.1% uncertainty). - Confirm knowledge of meter types (turbine, Coriolis, ultrasonic) and their operational principles. - Assess understanding of regulatory standards such as API MPMS Chapter 4 and its impact on fiscal reporting. - Ensure awareness of legal implications and audit requirements for custody transfer systems. |  |  |
| K. 3.1 | Describe the calibration process for custody transfer meters and associated transmitters. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 3.1 | - Verify understanding of zero/span checks for pressure, temperature, and density transmitters. - Confirm knowledge of meter proving techniques using bi-directional prover loops. - Assess ability to apply calibration tolerances and intervals based on process conditions. - Ensure awareness of traceability requirements for calibration results and compliance records. |  |  |
| K. 3.2 | Explain the integration of sampling systems with metering setups. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 3.2 | - Verify knowledge of flow-proportional and time-based sampling principles. - Confirm understanding of inline and bypass sampling configurations. - Assess ability to explain the role of sampling in validating metering accuracy. - Ensure understanding of international standards (e.g., ISO 3171, API MPMS Chapter 8). |  |  |
| K .3.3 | State the principles of BS&W (Basic Sediment and Water) monitoring and its calibration requirements. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 3.3 | - Verify knowledge of capacitance-based BS&W analyzer operation (e.g., Zelen Tech ZT-100). - Confirm understanding of calibration using oil/water reference standards. - Assess ability to identify errors caused by fouling, temperature, or flow changes. - Ensure awareness of API standards for water cut measurements (e.g., API MPMS Chapter 10.9). |  |  |
| K .3.4 | Describe troubleshooting techniques for sampling pumps, probes, and valves. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 3.4 | - Verify understanding of identifying flow restrictions, leaks, or signal loss. - Confirm knowledge of diagnostic techniques (e.g., pressure tests, flowrate checks). - Assess awareness of common faults such as pump cavitation or probe fouling. - Ensure understanding of restoring functionality through cleaning or component replacement. |  |  |
| K .3.5 | Explain the process for validating sample integrity for laboratory analysis. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 3.5 | - Verify understanding of contamination prevention during sample collection. - Confirm knowledge of ensuring representative samples using proportional sampling. - Assess awareness of proper storage, sealing, and labeling techniques for traceability. - Ensure understanding of matching lab results with inline analyzer data for validation. |  |  |
| K .3.6 | Explain compliance requirements for sampling and metering systems (API MPMS, ISO). |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 3.6 | - Verify understanding of API MPMS Chapters 4, 5, 6, 8, and ISO 3171 standards. - Confirm knowledge of reporting and audit requirements for metering and sampling systems. - Assess ability to ensure accuracy, repeatability, and traceability of measurements. - Ensure familiarity with third-party audit processes and government regulatory requirements. |  |  |
| K .3.7 | Describe the importance of flow conditioning and its impact on metering accuracy. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 3.7 | - Verify understanding of flow conditioners (e.g., straighteners, 10D/5D piping). - Confirm knowledge of mitigating turbulence to ensure laminar flow conditions. - Assess awareness of inaccuracies caused by improper flow profile or installation. - Ensure understanding of maintaining compliance with flowmeter manufacturer guidelines. |  |  |
| K .3.8 | Explain the process of real-time data integration with DCS/PLC systems. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 3.8 | - Verify knowledge of communication protocols (e.g., Modbus, HART, PROFIBUS). - Confirm understanding of integrating analyzer outputs with flow computers for real-time monitoring. - Assess awareness of data validation processes to ensure accuracy and reliability. - Ensure understanding of identifying and resolving interface faults between systems. |  |  |
| K .3.9 | Evaluate common failure modes in Gas Chromatographs. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 3.9 | - Detector drift due to contaminated sample or carrier gas impurities. - Column degradation from thermal stress or high molecular weight contaminants. - Signal loss caused by carrier gas pressure fluctuations or detector electronics failure. - Data transmission issues due to faulty RS-485/Modbus communication links. |  |  |
| K .3.10 | Assess regulatory compliance and safety requirements for Gas Chromatograph installations. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| KG. 3.10 | - Installation must comply with IEC 60079 for explosive atmosphere zones. - Analyzer cabinets must meet IP66 environmental protection ratings. - Calibration gases must be handled per NACE MR0175 standards. - System shutdown procedures must align with API RP 555 process analyzer guidelines. |  |  |
| Assessor Comments (provide a summary of the outcomes) |  |  |  |

| Assessor is to check the relevant methods used: 1. Observed 2. Simulated/Demonstrated 3. Questioned |  | 4. Professional Discussion 5. Witness Testimony 6. Other (Please comment on methods used) | Check Relevant Boxes |
| --- | --- | --- | --- |
| Code | Performance Criteria |  |  |
| Basic – Level 1 |  |  |  |
| P. 1.0 | Demonstrate understanding by verifying flowmeter accuracy and compliance with fiscal standards. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 1.0 | - Verify ability to check metering outputs against API MPMS Chapter 4 and 5 requirements. - Confirm the technician’s understanding of repeatability limits (e.g., ±0.01% during proving). - Assess ability to document flowrate accuracy, temperature, and density checks. - Ensure ability to escalate deviations in metering accuracy to supervisors. |  |  |
| P. 1.1 | Assist in operating and verifying the performance of the prover loop. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 1.1 | - Verify correct setup of the prover loop, including sphere launching and detector alignment. - Confirm ability to validate flowmeter outputs during proving cycles. - Assess ability to monitor prover sphere speed and pressure drop (e.g., 1.6 m/s max velocity). - Ensure proper documentation of proving results and verification of repeatability. |  |  |
| P. 1.2 | Verify DP transmitter accuracy and troubleshoot differential pressure issues. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 1.2 | - Confirm ability to check impulse tubing for leaks, blockages, or temperature effects. - Verify correct zero and span settings using HART communicators. - Assess ability to compare DP readings to expected flow outputs. - Ensure understanding of diagnosing signal loss caused by wiring or transmitter faults. |  |  |
| P. 1.3 | Perform configuration and diagnostics using HART communication tools. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 1.3 | - Verify ability to connect and access transmitter parameters using a HART communicator. - Confirm knowledge of adjusting zero/span values and reviewing primary variables (PV). - Assess ability to identify and clear fault codes or alarms through diagnostics. - Ensure proper verification of restored transmitter communication after configuration. |  |  |
| P. 1.4 | Inspect, verify, and troubleshoot flow measurement using a Heliflu-TZN turbine meter. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 1.4 | - Verify understanding of pulse outputs proportional to flow velocity. - Confirm ability to check meter bearings for wear or fouling. - Assess ability to monitor pressure drop across the turbine meter. - Ensure proper documentation of flow verification results. |  |  |
| P. 1.5 | Perform maintenance tasks on the Heliflu-TZN turbine meter. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 1.5 | - Verify ability to isolate and depressurize the meter before maintenance. - Confirm understanding of cleaning turbine blades to remove fouling or debris. - Assess ability to check for proper rotation and mechanical wear. - Ensure proper reinstallation and verification of turbine functionality. |  |  |
| P. 1.6 | Perform water cut verification using the Zelen Tech ZT-100 analyzer. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 1.6 | - Verify understanding of capacitance-based measurement and dielectric constant effects. - Confirm ability to verify analyzer outputs using manual grab samples. - Assess ability to check temperature compensation settings for accuracy. - Ensure correct adjustment of calibration values to match process requirements. |  |  |
| P. 1.7 | Diagnose and resolve faults with the ZT-100 water cut analyzer. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 1.7 | - Verify ability to identify sensor fouling, flowrate inconsistencies, and signal noise. - Confirm proper cleaning of sensor probes to resolve fouling-related faults. - Assess ability to recalibrate the analyzer after fault resolution. - Ensure accurate fault documentation and reporting to supervisors. |  |  |
| P. 1.8 | Verify viscosity readings and troubleshoot transmitter performance. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 1.8 | - Verify understanding of torsional oscillation and its role in viscosity measurement. - Confirm ability to test and compare viscosity readings against reference fluids. - Assess ability to identify temperature compensation issues impacting readings. - Ensure proper documentation of calibration verification results. |  |  |
| P. 1.9 | Perform maintenance tasks on the Emerson Micro Motion viscosity transmitter. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 1.9 | - Verify ability to isolate and depressurize the sensor assembly safely. - Confirm understanding of cleaning techniques to remove fouling or deposits. - Assess ability to inspect for physical damage or loss of sensitivity. - Ensure reinstallation and functionality testing post-maintenance. |  |  |
| P. 1.10 | Verify and configure data inputs on the ABB Spirit Flow-X flow computer. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 1.10 | - Verify ability to access flowrate, BS&W, and density data inputs for accuracy. - Confirm understanding of configuring alarms and thresholds. - Assess ability to validate communication with DP transmitters and analyzers. - Ensure proper documentation of configurations and diagnostic results. |  |  |
| P. 1.11 | Review and analyze alarm logs and diagnostics on the Flow-X system. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 1.11 | - Verify ability to access and interpret error logs for root cause analysis. - Confirm understanding of clearing alarms after resolving system faults. - Assess ability to escalate unresolved alarms with supporting diagnostics. - Ensure proper logging of alarm events and corrective actions. |  |  |
| P. 1.12 | Collect representative oil samples using the Auto Control CS55 grab sampler. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 1.12 | - Verify ability to flush sample points to ensure accuracy. - Confirm understanding of flow-proportional sample collection techniques. - Assess ability to label, store, and document samples correctly. - Ensure adherence to API MPMS Chapter 8 sampling guidelines. |  |  |
| P. 1.13 | Perform safety and operational checks on grab sampling equipment. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 1.13 | - Verify understanding of isolating and depressurizing sample valves before use. - Confirm ability to inspect seals and valves for leaks or wear. - Assess ability to check sample bottle cleanliness and readiness. - Ensure proper PPE use and adherence to safety protocols. |  |  |
| P. 1.14 | Compile and escalate metering discrepancy reports to supervisors. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 1.14 | - Verify ability to document discrepancies in readings, alarms, or diagnostic outputs. - Confirm understanding of including detailed data logs and supporting evidence. - Assess ability to escalate issues promptly with recommended corrective actions. - Ensure proper reporting format is followed for traceability and audit purposes. |  |  |
| P. 1.15 | Perform a visual inspection of Gas Chromatograph components. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 1.15 | - Check sample probe integrity and ensure correct installation depth. - Inspect carrier gas and calibration gas pressure regulators for leaks. - Verify status of sample flow indicators and pressure gauges. - Ensure analyzer cabinet doors are properly secured and labeled. |  |  |
| P. 1.16 | Conduct a manual calibration cycle on the Gas Chromatograph. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 1.16 | - Initiate calibration mode from local HMI or ICSS interface. - Introduce certified calibration gas into the sample stream. - Verify detector response and adjust peak retention times if required. - Document calibration results and report any deviations. |  |  |
| Intermediate – Level 2 |  |  |  |
| P. 2.0 | Supervise calibration and validation of water-in-oil analyzers (ZT-100). |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 2.0 | - Verify ability to coordinate calibration using certified reference standards. - Confirm understanding of validating analyzer outputs against manual samples. - Assess ability to troubleshoot probe fouling or signal noise. - Ensure accurate documentation of calibration results and adjustments. |  |  |
| P. 2.1 | Oversee the maintenance and functionality of automatic grab samplers (CS55). |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 2.1 | - Verify ability to coordinate routine maintenance such as valve cleaning and seal replacement. - Confirm understanding of flushing sampling points to avoid contamination. - Assess ability to validate correct flow-proportional sampling operation. - Ensure proper reporting of sampling equipment performance and issues. |  |  |
| P. 2.2 | Manage the calibration and performance of density analyzers (Micro Motion). |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 2.2 | - Verify ability to schedule and perform density verification using standard reference fluids. - Confirm ability to identify and address fouling or signal loss. - Assess ability to adjust transmitter parameters for process-specific conditions. - Ensure accurate logging of calibration data and fault diagnostics. |  |  |
| P. 2.3 | Oversee sampling campaigns and ensure compliance with ISO 3171/API MPMS. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 2.3 | - Verify ability to plan sampling activities aligned with production and regulatory requirements. - Confirm understanding of ensuring representative sample collection under varying flow conditions. - Assess ability to report sampling results and validate findings with laboratory analysis. - Ensure accurate documentation of sampling activities, findings, and recommendations. |  |  |
| P. 2.4 | Lead HSE-compliant maintenance activities on gas chromatographs and analyzers. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 2.4 | - Verify ability to enforce permit-to-work systems for hazardous area activities. - Confirm proper gas testing, isolation, and equipment depressurization. - Assess ability to monitor compliance with ATEX and IECEx standards. - Ensure technicians follow safe working practices, including PPE requirements. |  |  |
| P. 2.5 | Validate analyzer alarm logs and troubleshoot critical faults. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 2.5 | - Verify ability to review and prioritize critical alarms for immediate resolution. - Confirm understanding of fault diagnostics and identification of root causes. - Assess ability to coordinate corrective actions, including component replacement. - Ensure accurate reporting and escalation of unresolved faults. |  |  |
| P. 2.6 | Provide technical training to technicians on sampling systems and analyzers. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 2.6 | - Verify ability to develop competency-based training plans for sampling and analytical equipment. - Confirm understanding of mentoring technicians during calibration and troubleshooting. - Assess ability to provide feedback and evaluate technician performance. - Ensure maintenance of training records for audit and competency verification. |  |  |
| P. 2.7 | Audit documentation for calibration, maintenance, and sampling activities. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 2.7 | - Verify accuracy and completeness of calibration certificates and maintenance logs. - Confirm understanding of identifying and addressing discrepancies in records. - Assess ability to ensure traceability of all maintenance and sampling activities. - Ensure audit readiness by maintaining up-to-date compliance documentation. |  |  |
| P. 2.8 | Prepare performance reports for analyzers and sampling systems. |  | 1. ☐ 2. ☒ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 2.8 | - Verify ability to analyze trends, deviations, and system performance data. - Confirm understanding of compiling detailed reports for stakeholders. - Assess ability to recommend improvements based on report findings. - Ensure adherence to reporting timelines and standards for regulatory compliance. |  |  |
| P. 2.9 | Diagnose sample flow and pressure irregularities in the conditioning system. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 2.9 | - Measure sample line pressure drop using inline pressure sensors. - Inspect filtration units for blockages and saturation. - Check for leaks in stainless steel sample tubing and fittings. - Verify correct operation of bypass flowmeter and pressure relief valves. |  |  |
| P. 2.10 | Perform detector and column troubleshooting procedures. |  | 1. ☐ 2. ☒ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 2.10 | - Monitor baseline noise levels and peak resolution stability. - Replace or regenerate chromatographic columns showing peak broadening. - Verify detector response using standard gas mixtures. - Inspect TCD/FID circuits for power supply fluctuations. |  |  |
| Advanced – Level 3 |  |  |  |
| P. 3.0 | Perform calibration and verification of custody transfer meters. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 3.0 | - Verify ability to conduct zero/span checks and compare meter readings against certified standards. - Confirm accurate proving operations using bi-directional prover loops or master meters. - Assess ability to document calibration results and deviations for audit purposes. - Ensure proper adjustments to maintain accuracy within specified tolerances (e.g., ±0.1%). |  |  |
| P. 3.1 | Supervise and configure sampling systems for representative sample collection. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 3.1 | - Verify ability to adjust sampling rates based on flowmeter data. - Confirm proper setup of inline or bypass sampling systems to avoid bias. - Assess understanding of flushing procedures to ensure clean sampling points. - Ensure accurate documentation of sample conditions, including pressure, temperature, and flow. |  |  |
| P. 3.2 | Diagnose and resolve alarms or deviations in BS&W and density analyzers. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 3.2 | - Verify ability to analyze diagnostic logs for signal drift, fouling, or probe errors. - Confirm execution of recalibration using certified oil/water reference standards. - Assess ability to restore accuracy by inspecting and cleaning analyzer probes. - Ensure proper escalation of unresolved faults with supporting documentation. |  |  |
| P. 3.3 | Integrate real-time data from sampling and metering systems into DCS/PLC networks. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 3.3 | - Verify ability to troubleshoot communication failures and ensure stable analyzer outputs. - Confirm proper mapping of real-time data for custody transfer validation. - Assess ability to validate signal integrity and identify noise or data errors. - Ensure updates of firmware or software for flow computers and analyzers. |  |  |
| P. 3.4 | Perform system checks for flow conditioning devices impacting metering accuracy. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 3.4 | - Verify ability to inspect straight runs and flow conditioners for installation compliance. - Confirm understanding of pressure drop checks across flow conditioning elements. - Assess ability to troubleshoot flow anomalies caused by upstream turbulence. - Ensure accurate reporting and resolution of identified issues. |  |  |
| P. 3.5 | Validate sample integrity through comparison with lab results. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 3. 5 | - Verify ability to reconcile inline analyzer data with laboratory analysis. - Confirm proper interpretation of discrepancies in BS&W or density results. - Assess ability to escalate unresolvable variances for detailed investigation. - Ensure thorough documentation of validation activities. |  |  |
| P. 3.6 | Support Factory Acceptance Tests (FAT) and Site Acceptance Tests (SAT) for metering equipment. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 3. 6 | - Verify ability to review FAT/SAT procedures and test specifications. - Confirm hands-on validation of equipment performance against acceptance criteria. - Assess understanding of documenting and reporting deviations observed during testing. - Ensure recommendations for commissioning and post-installation performance checks. |  |  |
| P. 3.7 | Optimize Gas Chromatograph performance using diagnostic tools. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 3. 7 | - Analyze chromatograms for signal noise, retention time shifts, and peak resolution. - Adjust column oven temperature and carrier gas flow rate for process variations. - Use SCADA interface and RS-485 diagnostics to detect communication faults. - Implement predictive maintenance schedules for component replacement. |  |  |
| P. 3.8 | Lead a root cause failure analysis (RCFA) for chronic Gas Chromatograph faults. |  | 1. ☐ 2. ☐ 3. ☐ 4. ☐ 5. ☐ 6. ☐ |
| PG. 3. 8 | - Collect historical alarm logs and trending data. - Conduct failure mode analysis on detector, carrier gas system, and control electronics. - Evaluate sample system contamination risks and recommend mitigation strategies. - Develop corrective action plans based on API RP 555 process analyzer guidelines. |  |  |
| Assessor Comments (provide a summary of the outcomes) |  |  |  |

| Additional Assessment Evidence & References (Please reference, scan, and upload the evidence to candidate’s portfolio along with this report) |  |  |
| --- | --- | --- |
| Ref. No. | Type | Title |